## Profile
| Device Profile: | 5001 | Modular Device Profile |  | Usage Notes: - The PDO mapping object and SyncManager assignment object doesn't need to be defined. In that case they are created automatically. - The following objects are fixed included in the SSC and shall not be defined in the file : 0x1000, 0x1001, 0x1008, 0x1009, 0x100a, 0x1010, 0x1011, 0x1018, 0x10F0, 0x10F1, 0x10F3, 0x1c00, 0x1c32, 0x1c33 - Entries less or equal one 8Bit shall not overlap byte borders. - Entries greater 8Bit shall always start at an exact word border
| Modul Profile: | 0
The object dictionary defined here shall be used complementary with ETG.5001 and ETG.1000
| Index | ObjectCode | SI | DataType | Name | Default | Min | Max | M/O/C | B/S | Access | rx/tx | CoeRead | CoeWrite | Description
| //0x04xx | Units
| //0x8xx | Enums (0x0800 - 0xFFF)
| //0x14nn | RxPDO Parameter (0x1400 - 0x15FF)
| //0x16nn | RxPDO Mapping (0x1600 - 0x17FF) ¦ NOTE: if no RxPDO mapping object is defined the will be created automatically
| //0x18nn | TxPDO Paramter (0x1800 - 0x19FF)
| //0x1Ann | TxPDO Mapping (0x1A00 - 0x1BFF) ¦ NOTE: if no TxPDO mapping object is defined the will be created automatically
| //0x1C12 | SyncManager 2 Assignment ¦ NOTE: if this object is not defined it will be created automatically
| //0x1C13 | SyncManager 3 Assignment ¦ NOTE: if this object is not defined it will be created automatically
| //0x6nnx | Input Data of the Module (0x6000 - 0x6FFF)
| 0x6000 | RECORD |  |  | InputData
|  |  | 1 | UNSIGNED8 | InU8_BoardID_RTM_ON_ARM | [9 empty] | bit0~2:单板硬件设计版本，与设计保持一致，01开始 bit3~7:每个单板的唯一硬件ID，从0起始；
|  |  | 2 | UNSIGNED8 | InU8_Reserved_RTM_ON_ARM
|  |  | 3 | UNSIGNED32 | InU32_FirmWareVersion_RTM_ON_ARM | [9 empty] | 固件版本，"00.00.00.01"
|  |  | 4 | UNSIGNED8 | InU8_BoardID_ICM
|  |  | 5 | UNSIGNED8 | InU8_Reserved_ICM
|  |  | 6 | UNSIGNED32 | InU32_FirmWareVersion_ICM
|  |  | 7 | UNSIGNED8 | InU8_BoardID_BGM
|  |  | 8 | UNSIGNED8 | InU8_Reserved_BGM
|  |  | 9 | UNSIGNED32 | InU32_FirmWareVersion_BGM
|  |  | 10 | UNSIGNED8 | InU8_BoardID_primary_dose
|  |  | 11 | UNSIGNED8 | InU8_Reserved_primary_dose
|  |  | 12 | UNSIGNED32 | InU32_FirmWareVersion_primary_dose
|  |  | 13 | UNSIGNED8 | InU8_BoardID_second_dose
|  |  | 14 | UNSIGNED8 | InU8_Reserved_second_dose
|  |  | 15 | UNSIGNED32 | InU32_FirmWareVersion_second_dose
|  |  | 16 | UNSIGNED8 | InU8_BoardID_AFC
|  |  | 17 | UNSIGNED8 | InU8_Reserved_AFC
|  |  | 18 | UNSIGNED32 | InU32_FirmWareVersion_AFC
|  |  | 19 | UNSIGNED8 | InU8_BoardID_QAM
|  |  | 20 | UNSIGNED8 | InU8_Reserved_QAM
|  |  | 21 | UNSIGNED32 | InU32_FirmWareVersion_QAM
|  |  | 22 | UNSIGNED8 | InU8_BoardID_RTM_OFF_ARM
|  |  | 23 | UNSIGNED8 | InU8_Reserved_RTM_OFF_ARM
|  |  | 24 | UNSIGNED32 | InU32_FirmWareVersion_RTM_OFF_ARM
|  |  | 25 | UNSIGNED8 | InU8_BoardID_PSM
|  |  | 26 | UNSIGNED8 | InU8_Reserved_PSM
|  |  | 27 | UNSIGNED32 | InU32_FirmWareVersion_PSM
|  |  | 28 | UNSIGNED8 | InU8_BoardID_GMM
|  |  | 29 | UNSIGNED8 | InU8_Reserved_GMM
|  |  | 30 | UNSIGNED32 | InU32_FirmWareVersion_GMM
|  |  | 31 | UNSIGNED8 | InU8_BoardID_FKP
|  |  | 32 | UNSIGNED8 | InU8_Reserved_FKP
|  |  | 33 | UNSIGNED32 | InU32_FirmWareVersion_FKP
|  |  | 34 | UNSIGNED8 | InU8_BoardID_CPG_L
|  |  | 35 | UNSIGNED8 | InU8_Reserved_CPG_L
|  |  | 36 | UNSIGNED32 | InU32_FirmWareVersion_CPG_L
|  |  | 37 | UNSIGNED8 | InU8_BoardID_CPG_R
|  |  | 38 | UNSIGNED8 | InU8_Reserved_CPG_R
|  |  | 39 | UNSIGNED32 | InU32_FirmWareVersion_CPG_R
|  |  | 40 | UNSIGNED8 | InU8_beam_id
4 |  |  | 41 | UNSIGNED8 | InU8_reserved1
|  |  | 42 | UINT | InU16_radiation_index
|  |  | 43 | UNSIGNED8 | InU8_rtm_on_arm_fsm_state_current
|  |  | 44 | UNSIGNED8 | InU8_rtm_on_arm_ctrl_mode_cur
|  |  | 45 | UNSIGNED8 | InU8_ethercat_Link_state
|  |  | 46 | UNSIGNED8 | InU8_reserved2
|  |  | 47 | UNSIGNED8 | InU8_rtm_on_arm_beam_id
|  |  | 48 | UNSIGNED8 | InU8_reserved3
|  |  | 49 | UINT | InU16_rtm_on_arm_radiation_index
|  |  | 50 | UNSIGNED32 | InU32_rtm_on_arm_not_ready_event
|  |  | 51 | UNSIGNED32 | InU32_rtm_on_arm_warning_interlock
|  |  | 52 | UNSIGNED32 | InU32_rtm_on_arm_minor_interlock
|  |  | 53 | UNSIGNED32 | InU32_rtm_on_arm_serious_interlock
|  |  | 54 | UNSIGNED16 | InU16_TotalStep_RTM_ON_ARM
|  |  | 55 | UNSIGNED16 | InU16_CurrentStep_RTM_ON_ARM
|  |  | 56 | UNSIGNED16 | InU16_ErrorCode_RTM_ON_ARM
|  |  | 57 | BOOLEAN | InB1_ON_DI_RTC_WD_OK_IN
|  |  | 58 | BOOLEAN | InB1_ON_DI_Reserve0
|  |  | 59 | BOOLEAN | InB1_ON_DI_BSM_NOT_READY
|  |  | 60 | BOOLEAN | InB1_ON_DI_MV_TreatmentEN
|  |  | 61 | BOOLEAN | InB1_ON_DI_DI_HVEN
|  |  | 62 | BOOLEAN | InB1_ON_DI_Pulse_Inhibit
|  |  | 63 | BOOLEAN | InB1_ON_DI_KV_TreatmentEN
|  |  | 64 | BOOLEAN | InB1_ON_DI_Power_cut_FB
|  |  | 65 | BOOLEAN | InB1_ON_DI_GATING_IN
|  |  | 66 | BOOLEAN | InB1_ON_DI_Slipring_HVEN_IN
|  |  | 67 | BOOLEAN | InB1_ON_DI_Slipring_KV_TreatmentEN_IN
|  |  | 68 | BOOLEAN | InB1_ON_DI_Slipring_MV_TreatmentEN_IN
|  |  | 69 | BIT4 | InB4_ON_DI_reserve1
|  |  | 70 | BOOLEAN | InB1_ON_DO_reserve0
|  |  | 71 | BOOLEAN | InB1_ON_DO_MV_TreatmentEN
|  |  | 72 | BOOLEAN | InB1_ON_DO_KV_TreatmentEN
|  |  | 73 | BOOLEAN | InB1_ON_DO_Emergency
|  |  | 74 | BOOLEAN | InB1_ON_DO_Pulse_Inhibit
|  |  | 75 | BOOLEAN | InB1_ON_DO_HVEN
|  |  | 76 | BOOLEAN | InB1_ON_DO_PowerCut
|  |  | 77 | BOOLEAN | InB1_ON_DO_reserve1
|  |  | 78 | UNSIGNED8 | InU8_ON_DO_reserve2
|  |  | 79 | UNSIGNED8 | InU8_OFF_reserve0
|  |  | 80 | BOOLEAN | InB1_OFF_DI_CITB_EMERGENCY2
|  |  | 81 | BOOLEAN | InB1_OFF_DI_UPS_LOAD_PORT
|  |  | 82 | BOOLEAN | InB1_OFF_DI_STAND_BREAKER1
|  |  | 83 | BOOLEAN | InB1_OFF_DI_STAND_EMERGENCY
|  |  | 84 | BOOLEAN | InB1_OFF_DI_TouchGuard
|  |  | 85 | BOOLEAN | InB1_OFF_DI_HvEn
|  |  | 86 | BOOLEAN | InB1_OFF_DI_MV_TreatmentEN
|  |  | 87 | BOOLEAN | InB1_OFF_DI_KV_TreatmentEN
|  |  | 88 | BOOLEAN | InB1_OFF_DI_STAND_BREAKER7
|  |  | 89 | BOOLEAN | InB1_OFF_DI_STAND_BREAKER3
|  |  | 90 | BOOLEAN | InB1_OFF_DI_UPS_ON_BYPASS
|  |  | 91 | BOOLEAN | InB1_OFF_DI_STAND_BREAKER5
|  |  | 92 | BOOLEAN | InB1_OFF_DI_CITB_EMERGENCY3
|  |  | 93 | BOOLEAN | InB1_OFF_DI_COVER3
|  |  | 94 | BOOLEAN | InB1_OFF_DI_POWER_CUT
|  |  | 95 | BOOLEAN | InB1_OFF_DI_STAND_BREAKER4
|  |  | 96 | BOOLEAN | InB1_OFF_DI_CITB_TREATMENT_ROOM_DOOR1
|  |  | 97 | BOOLEAN | InB1_OFF_DI_CITB_TREATMENT_ROOM_DOOR2
|  |  | 98 | BOOLEAN | InB1_OFF_DI_CITB_EMERGENCY1
|  |  | 99 | BOOLEAN | InB1_OFF_DI_CITB_SEARCH_TREATMENT_ROOM
|  |  | 100 | BOOLEAN | InB1_OFF_DI_CONTROL_ROOM_BREAKER1
|  |  | 101 | BOOLEAN | InB1_OFF_DI_UPS_ON_BATT
|  |  | 102 | BOOLEAN | InB1_OFF_DI_CONTROL_ROOM_BREAKER5
|  |  | 103 | BOOLEAN | InB1_OFF_DI_STAND_BREAKER6
|  |  | 104 | BOOLEAN | InB1_OFF_DI_CITB_EMERGENCY5
|  |  | 105 | BOOLEAN | InB1_OFF_DI_COVER1
|  |  | 106 | BOOLEAN | InB1_OFF_DI_STAND_RESERVE
|  |  | 107 | BOOLEAN | InB1_OFF_DI_TREATMENT_ROOM_DOOR_READY
|  |  | 108 | BOOLEAN | InB1_OFF_DI_COVER2
|  |  | 109 | BOOLEAN | InB1_OFF_DI_CONTROL_ROOM_BREAKER2
|  |  | 110 | BOOLEAN | InB1_OFF_DI_CONTROL_ROOM_BREAKER4
|  |  | 111 | BOOLEAN | InB1_OFF_DI_STAND_CONTACTOR2
|  |  | 112 | BOOLEAN | InB1_OFF_DI_STAND_CONTACTOR1
|  |  | 113 | BOOLEAN | InB1_OFF_DI_UPS_LOW_BATT
|  |  | 114 | BOOLEAN | InB1_OFF_DI_HvKey
|  |  | 115 | BOOLEAN | InB1_OFF_DI_COVER4
|  |  | 116 | BOOLEAN | InB1_OFF_DI_CITB_EMERGENCY4
|  |  | 117 | BOOLEAN | InB1_OFF_DI_CONTROL_ROOM_BREAKER3
|  |  | 118 | BOOLEAN | InB1_OFF_DI_STAND_BREAKER2
|  |  | 119 | BOOLEAN | InB1_OFF_RTC_WD_OK_IN
|  |  | 120 | UNSIGNED16 | InU16_OFF_mcp23017_di_reserve
|  |  | 121 | BOOLEAN | InB1_OFF_DI_GATING
|  |  | 122 | BIT7 | InB7_OFF_reserve1
|  |  | 123 | UNSIGNED8 | InU8_OFF_reserve2
|  |  | 124 | BOOLEAN | InB1_OFF_DO_RTM_AutoPowerUp
|  |  | 125 | BOOLEAN | InB1_OFF_DO_SearchTreatmentRoomRelay
|  |  | 126 | BOOLEAN | InB1_OFF_DO_Laser
|  |  | 127 | BOOLEAN | InB1_OFF_DO_Power_CUT
|  |  | 128 | BOOLEAN | InB1_OFF_DO_STAND_RESERVE
|  |  | 129 | BOOLEAN | InB1_OFF_DO_TreatmentRoomLight
|  |  | 130 | BOOLEAN | InB1_OFF_DO_RadiationIndicator
|  |  | 131 | BOOLEAN | InB1_OFF_DO_ReadyIndicator
|  |  | 132 | BOOLEAN | InB1_OFF_DO_RTM_SystemShutDown
|  |  | 133 | BIT7 | InB7_OFF_reserve3
|  |  | 134 | UNSIGNED16 | InU16_OFF_mcp23017_do_reserve
|  |  | 135 | BOOLEAN | InB1_OFF_DO_SoftwareHvEn
|  |  | 136 | BOOLEAN | InB1_OFF_DO_SoftwareKVTreatmentEn
|  |  | 137 | BOOLEAN | InB1_OFF_DO_SoftwareMVTreatmentEn
|  |  | 138 | BOOLEAN | InB1_OFF_DO_ThreePhasePowerOn
|  |  | 139 | BOOLEAN | InB1_OFF_DO_softwareMoveEN
|  |  | 140 | BOOLEAN | InB1_OFF_DO_TreatmentMotionEnable
|  |  | 141 | BOOLEAN | InB1_OFF_DO_AsuMotionEnable
|  |  | 142 | BOOLEAN | InB1_OFF_reserve4
|  |  | 143 | UNSIGNED8 | InU8_OFF_reserve5
|  |  | 144 | UNSIGNED8 | InU8_icm_fsm_state_current
|  |  | 145 | UNSIGNED8 | InU8_icm_ctrl_mode_cur
|  |  | 146 | UNSIGNED16 | InU16_reserved4
|  |  | 147 | UNSIGNED8 | InU8_icm_beam_id
|  |  | 148 | UNSIGNED8 | InU8_reserved5
|  |  | 149 | UINT | InU16_icm_radiation_index
|  |  | 150 | UNSIGNED32 | InU32_icm_not_ready_event
|  |  | 151 | UNSIGNED32 | InU32_icm_warning_interlock
|  |  | 152 | UNSIGNED32 | InU32_icm_minor_interlock
|  |  | 153 | UNSIGNED32 | InU32_icm_serious_interlock
|  |  | 154 | UNSIGNED16 | InU16_TotalStep_ICM
|  |  | 155 | UNSIGNED16 | InU16_CurrentStep_ICM
|  |  | 156 | UNSIGNED16 | InU16_ErrorCode_ICM
|  |  | 157 | UNSIGNED8 | InU8_bgm_fsm_state_current
|  |  | 158 | UNSIGNED8 | InU8_bgm_ctrl_mode_cur
|  |  | 159 | UNSIGNED16 | InU16_reserved6
|  |  | 160 | UNSIGNED8 | InU8_bgm_beam_id
|  |  | 161 | UNSIGNED8 | InU8_reserved7
|  |  | 162 | UINT | InU16_bgm_radiation_index
|  |  | 163 | UNSIGNED32 | InU32_bgm_not_ready_event
|  |  | 164 | UNSIGNED32 | InU32_bgm_warning_interlock
|  |  | 165 | UNSIGNED32 | InU32_bgm_minor_interlock
|  |  | 166 | UNSIGNED32 | InU32_bgm_serious_interlock
|  |  | 167 | UNSIGNED16 | InU16_TotalStep_BGM
|  |  | 168 | UNSIGNED16 | InU16_CurrentStep_BGM
|  |  | 169 | UNSIGNED16 | InU16_ErrorCode_BGM
|  |  | 170 | REAL | InF_beam_on_time
|  |  | 171 | REAL | InF_primary_dose_current
|  |  | 172 | REAL | InF_primary_dose_rate_current
|  |  | 173 | UNSIGNED8 | InU8_primary_dose_fsm_state
|  |  | 174 | UNSIGNED8 | InU8_reserved8
|  |  | 175 | UNSIGNED32 | InU32_primary_dose_interlock
|  |  | 176 | REAL | InF_second_dose_current
|  |  | 177 | REAL | InF_secondry_dose_rate_current
|  |  | 178 | UNSIGNED8 | InU8_secondry_dose_fsm_rate
|  |  | 179 | UNSIGNED8 | InU8_reserved9
|  |  | 180 | UNSIGNED32 | InU32_secondry_dose_interlock
|  |  | 181 | UNSIGNED32 | InU32_afc_interlock
|  |  | 182 | REAL | InF_afc_position_current
|  |  | 183 | UNSIGNED8 | InU8_qam_fsm_state_current
|  |  | 184 | UNSIGNED8 | InU8_qam_ctrl_mode_cur
|  |  | 185 | UNSIGNED16 | InU16_reserved10
|  |  | 186 | UNSIGNED8 | InU8_qam_beam_id
|  |  | 187 | UNSIGNED8 | InU8_reserved11
|  |  | 188 | UINT | InU16_qam_radiation_index
|  |  | 189 | UNSIGNED32 | InU32_qam_not_ready_event
|  |  | 190 | UNSIGNED32 | InU32_qam_warning_interlock
|  |  | 191 | UNSIGNED32 | InU32_qam_minor_interlock
|  |  | 192 | UNSIGNED32 | InU32_qam_serious_interlock
|  |  | 193 | UNSIGNED16 | InU16_TotalStep_QAM
|  |  | 194 | UNSIGNED16 | InU16_CurrentStep_QAM
|  |  | 195 | UNSIGNED16 | InU16_ErrorCode_QAM
|  |  | 196 | UNSIGNED8 | InU8_rtm_off_fsm_state_current
|  |  | 197 | UNSIGNED8 | InU8_rtm_off_ctrl_mode_cur
|  |  | 198 | UNSIGNED16 | InU16_reserved12
|  |  | 199 | UNSIGNED8 | InU8_rtm_off_beam_id
|  |  | 200 | UNSIGNED8 | InU8_reserved13
|  |  | 201 | UINT | InU16_rtm_off_radiation_index
|  |  | 202 | UNSIGNED32 | InU32_rtm_off_not_ready_event
|  |  | 203 | UNSIGNED32 | InU32_rtm_off_warning_interlock
|  |  | 204 | UNSIGNED32 | InU32_rtm_off_minor_interlock
|  |  | 205 | UNSIGNED32 | InU32_rtm_off_serious_interlock
|  |  | 206 | UNSIGNED16 | InU16_TotalStep_RTM_OFF_ARM
|  |  | 207 | UNSIGNED16 | InU16_CurrentStep_RTM_OFF_ARM
|  |  | 208 | UNSIGNED16 | InU16_ErrorCode_RTM_OFF_ARM
|  |  | 209 | UNSIGNED32 | InU32_rtm_off_search_timeout_cur
|  |  | 210 | UNSIGNED8 | InU8_gmm_fsm_state_current
|  |  | 211 | UNSIGNED8 | InU8_gmm_ctrl_mode_cur
|  |  | 212 | UNSIGNED16 | InU16_reserved14
|  |  | 213 | UNSIGNED8 | InU8_gmm_beam_id
|  |  | 214 | UNSIGNED8 | InU8_reserved15
|  |  | 215 | UINT | InU16_gmm_radiation_index
|  |  | 216 | UNSIGNED32 | InU32_gmm_not_ready_event
|  |  | 217 | UNSIGNED32 | InU32_gmm_warning_interlock
|  |  | 218 | UNSIGNED32 | InU32_gmm_minor_interlock
|  |  | 219 | UNSIGNED32 | InU32_gmm_serious_interlock
|  |  | 220 | UNSIGNED16 | InU16_TotalStep_GMM
|  |  | 221 | UNSIGNED16 | InU16_CurrentStep_GMM
|  |  | 222 | UNSIGNED16 | InU16_ErrorCode_GMM
|  |  | 223 | UNSIGNED32 | InU32_gmm_move_status
|  |  | 224 | REAL | InF_gmm_position_cur
|  |  | 225 | REAL | InF_gmm_velocity_cur
|  |  | 226 | UNSIGNED8 | InU8_psm_fsm_state_current
|  |  | 227 | UNSIGNED8 | InU8_psm_ctrl_mode_cur
|  |  | 228 | UNSIGNED16 | InU16_reserved16
|  |  | 229 | UNSIGNED8 | InU8_psm_beam_id
|  |  | 230 | UNSIGNED8 | InU8_reserved17
|  |  | 231 | UINT | InU16_psm_radiation_index
|  |  | 232 | UNSIGNED32 | InU32_psm_not_ready_event
|  |  | 233 | UNSIGNED32 | InU32_psm_warning_interlock
|  |  | 234 | UNSIGNED32 | InU32_psm_minor_interlock
|  |  | 235 | UNSIGNED32 | InU32_psm_serious_interlock
|  |  | 236 | UNSIGNED16 | InU16_TotalStep_PSM
|  |  | 237 | UNSIGNED16 | InU16_CurrentStep_PSM
|  |  | 238 | UNSIGNED16 | InU16_ErrorCode_PSM
|  |  | 239 | UNSIGNED32 | InU32_psm_move_status
|  |  | 240 | REAL | InF_psm_position_x_cur
|  |  | 241 | REAL | InF_psm_position_y_cur
|  |  | 242 | REAL | InF_psm_position_z_cur
|  |  | 243 | REAL | InF_psm_position_x_r_cur
|  |  | 244 | REAL | InF_psm_position_y_r_cur
|  |  | 245 | REAL | InF_psm_position_z_r_cur
|  |  | 246 | REAL | InF_psm_velocity_x_cur
|  |  | 247 | REAL | InF_psm_velocity_y_cur
|  |  | 248 | REAL | InF_psm_velocity_z_cur
|  |  | 249 | REAL | InF_psm_velocity_x_r_cur
|  |  | 250 | REAL | InF_psm_velocity_y_r_cur
|  |  | 251 | REAL | InF_psm_velocity_z_r_cur
|  |  | 252 | UNSIGNED8 | InU8_function_button
|  |  | 253 | UNSIGNED8 | InU8_special_button
| 0x6010 | RECORD |  |  | InputData
|  |  | 1 | UNSIGNED32 | InU32_FkpButtonFaultMask
|  |  | 2 | UNSIGNED16 | InU16_TotalStep_FKP
|  |  | 3 | UNSIGNED16 | InU16_CurrentStep_FKP
|  |  | 4 | UNSIGNED16 | InU16_ErrorCode_FKP
|  |  | 5 | UNSIGNED32 | InU32_CpgButtonFaultMask_L
|  |  | 6 | UNSIGNED16 | InU16_TotalStep_CPG_L
|  |  | 7 | UNSIGNED16 | InU16_CurrentStep_CPG_L
|  |  | 8 | UNSIGNED16 | InU16_ErrorCode_CPG_L
|  |  | 9 | UNSIGNED32 | InU32_CpgButtonFaultMask_R
|  |  | 10 | UNSIGNED16 | InU16_TotalStep_CPG_R
|  |  | 11 | UNSIGNED16 | InU16_CurrentStep_CPG_R
|  |  | 12 | UNSIGNED16 | InU16_ErrorCode_CPG_R
|  |  | 13 | UNSIGNED16 | InU16_data_valid_flag
|  |  | 14 | REAL | InUF_load_position_x_cur
|  |  | 15 | REAL | InUF_load_position_y_cur
|  |  | 16 | REAL | InUF_load_position_z_cur
|  |  | 17 | REAL | InUF_load_position_x_r_cur
|  |  | 18 | REAL | InUF_load_position_y_r_cur
|  |  | 19 | REAL | InUF_load_position_z_r_cur
|  |  | 20 | REAL | InUF_position_x_tar
|  |  | 21 | REAL | InUF_position_y_tar
|  |  | 22 | REAL | InUF_position_z_tar
|  |  | 23 | REAL | InUF_position_x_r_tar
|  |  | 24 | REAL | InUF_position_y_r_tar
|  |  | 25 | REAL | InUF_position_z_r_tar
|  |  | 26 | UNSIGNED8 | InU8_trm_require_state
| //0x7nnx | Output Data of the Module (0x7000 - 0x7FFF)
| 0x7010 | RECORD |  |  | OutputData
|  |  | 1 | UNSIGNED8 | OutU8_BoardID
|  |  | 2 | UNSIGNED8 | OutU8_Reserved0
|  |  | 3 | UNSIGNED32 | OutU32_FirmWareVersion
|  |  | 4 | UNSIGNED8 | OutU8_beam_id
|  |  | 5 | UNSIGNED8 | OutU8_state_sync
|  |  | 6 | UINT | OutU16_radiation_index
|  |  | 7 | UNSIGNED8 | OutU8_fault_clear
|  |  | 8 | UNSIGNED8 | OutU8_systemCurrentState
|  |  | 9 | UINT | OutU16_reboot
|  |  | 10 | UNSIGNED8 | OutU8_treatment_record_state
|  |  | 11 | UNSIGNED8 | OutU8_ethercat_Link_state
|  |  | 12 | UNSIGNED8 | OutU8_rtm_on_require_state
|  |  | 13 | UNSIGNED8 | OutU8_rtm_on_require_ctrl_mode
|  |  | 14 | UNSIGNED16 | OutU16_rtm_on_plc_info
|  |  | 15 | UNSIGNED32 | OutU32_rtm_on_interlock_override
|  |  | 16 | UNSIGNED32 | OutU32_rtm_on_unready_override
|  |  | 17 | UNSIGNED16 | OutU16_rtm_on_treatment_mode
|  |  | 18 | UNSIGNED8 | OutU8_icm_require_state
|  |  | 19 | UNSIGNED8 | OutU8_icm_require_ctrl_mode
|  |  | 20 | UINT | OutU16_ct_status
|  |  | 21 | REAL | OutF_gantry_velocity_prepare
|  |  | 22 | UNSIGNED32 | OutU32_icm_interlock_override
|  |  | 23 | UNSIGNED32 | OutU32_icm_unready_override
|  |  | 24 | UNSIGNED8 | OutU8_bgm_require_state
|  |  | 25 | UNSIGNED8 | OutU8_bgm_require_ctrl_mode
|  |  | 26 | UNSIGNED32 | OutU32_bgm_interlock_override
|  |  | 27 | UNSIGNED32 | OutU32_bgm_unready_override
|  |  | 28 | UNSIGNED8 | OutU8_qam_require_state
|  |  | 29 | UNSIGNED8 | OutU8_qam_require_ctrl_mode
|  |  | 30 | UNSIGNED32 | OutU32_qam_interlock_override
|  |  | 31 | UNSIGNED32 | OutU32_qam_unready_override
|  |  | 32 | UNSIGNED8 | OutU8_bsm_require_state
|  |  | 33 | UNSIGNED8 | OutU8_bsm_require_ctrl_mode
|  |  | 34 | UNSIGNED32 | OutU32_bsm_interlock_override
|  |  | 35 | UNSIGNED32 | OutU32_bsm_unready_override
|  |  | 36 | UNSIGNED8 | OutU8_rtm_off_require_state
|  |  | 37 | UNSIGNED8 | OutU8_rtm_off_require_ctrl_mode
|  |  | 38 | UNSIGNED16 | OutU16_rtm_off_plc_info
|  |  | 39 | UNSIGNED32 | OutU32_rtm_off_interlock_override
|  |  | 40 | UNSIGNED32 | OutU32_rtm_off_unready_override
|  |  | 41 | UNSIGNED8 | OutU8_led_belt
|  |  | 42 | UNSIGNED8 | OutU8_reserved2
|  |  | 43 | UNSIGNED32 | OutU32_rtm_off_search_timeout
|  |  | 44 | UNSIGNED16 | OutU16_rtm_off_treatment_mode
|  |  | 45 | UNSIGNED8 | OutU8_gmm_require_state
|  |  | 46 | UNSIGNED8 | OutU8_gmm_require_ctrl_mode
|  |  | 47 | UNSIGNED32 | OutU32_gmm_interlock_override
|  |  | 48 | UNSIGNED32 | OutU32_gmm_unready_override
|  |  | 49 | UNSIGNED8 | OutU8_psm_require_state
|  |  | 50 | UNSIGNED8 | OutU8_psm_require_ctrl_mode
|  |  | 51 | UNSIGNED32 | OutU32_psm_interlock_override
|  |  | 52 | UNSIGNED32 | OutU32_psm_unready_override
|  |  | 53 | UNSIGNED16 | OutU16_fkp_year
|  |  | 54 | UNSIGNED8 | OutU8_fkp_month
|  |  | 55 | UNSIGNED8 | OutU8_fkp_day
|  |  | 56 | UNSIGNED8 | OutU8_fkp_hour
|  |  | 57 | UNSIGNED8 | OutU8_fkp_minute
|  |  | 58 | UNSIGNED8 | OutU8_fkp_totalFractions
|  |  | 59 | UNSIGNED8 | OutU8_fkp_fractions
|  |  | 60 | UNSIGNED8 | OutU8_reserved3
| //0x8nnx | Configuration Data of the Module (0x8000 - 0x8FFF)
| //0x9nnx | Information Data of the Module (0x9000 - 0x9FFF)
| //0xAnnx | Diagnosis Data of the Module (0xA000 - 0xAFFF)